## Sheet1
Unsafe Hot Plates*
Brand | Model | Spontaneous Heating | Runaway Heating | Other/Notes
Corning | PC-35 | X |  | No temperature feedback.
| PC-351 | X |  | No temperature feedback.
| PC-200 |  |  | "Off" does not disconnect power from heating element.
| PC-220 |  | X | "Off" does not disconnect power from heating element.
| PC-320 | X
| PC-400D | X
| PC-420 |  | X | "Off" does not disconnect power from heating element.
| PC-420D | X | X | "Off" does not disconnect power from heating element.
Troemner | 97042-714 Professional 97042-642 Advanced |  |  | MIT electronics tech evaluated 2 models of Troemner hot plates and found safety hazards based on the design and manufacturing practices.
Fisher | Isotemp 11-600-49H/ 11-700-49H |  |  | "Off" does not disconnect power from heating element.
Thermo/Fisher recalled certain Isotemp hot plates manufactured in 2015 due to defective on/off knobs
Chemglas | Optimag-St CG-1994-10 & -50/ CG-1993-T-50 | X |  | MIT electronics tech verified that this model heated to 300 within a min even though it was off. Refer to note about Chemglas on 2nd page.
Cimarec** | SP46925 | X |  | No temperature feedback.
| H-4954.xx |  | X
Thermo/Fisher recalled certain Cimarec hot plates manufactured in 2015 due to defective on/off knobs
** Sold under Thermolyne, Barnstead/Thermolyne, and Fisher names depending on age.
Notes
Exposure to elements such as cold, moisture, corrosive gas and vapor are known to increase the chance of processor failures, although problems have also been observed in nearly-new equipment.
Hot plates manufactured prior to 1984 do not have temperature feedback controls and can spontaneously and rapidly heat beyond the set temperature while in the ‘ON’ position.
MIT EHS recommends that you only use hot plates with the following safety features • Listed by UL or other Nationally Recognized Testing Lab (OSHA requirement) • Separate on/off switch from the heat/stir controls • Automatic “over-temperature” monitoring including if the external probe falls out (fail safe) • Hot surface light
Refer to webpage for additional details - https://ehs.mit.edu/site/laboratory-safety/unsafe-hot-plates.
The following brands have sold hot plates that do not have all of the safety features listed above or that have other problematic issues.
IKA | Model 2753 | Not NRTL listed as of 3/2018.
| CMAG | Regular CMAG is UL listed and has a fixed safety circuit of 550. The plate will shut off if temp goes over 500.
| CMAG - Digital version | Digital CMAG is UL listed and has adjustable safety circuit. CMAG HS-7 digital can error out easily and has a tendency to overshoot when heating.
| RCT B S001 | This model is UL listed, has a safety temperature limit, a Hot Top indicator, and an on/off switch.
Chemglass | CG-1997-V 35mm round top DIGITAL | Webpage implies that this hot plate is UL listed, but it is actually Nemko listed. Nemko is a nationally recognized testing labs in USA. This model has a Separate Programmable Maximum Safety Temperature and a Hot Top indicator.
Sources*
https://dchas.org/wp-content/uploads/2017/09/RunawayHotPlates-2017Aug.pdf
http://www.ehrs.upenn.edu/programs/labsafety/alerts/hot_plate_malfunction.html
http://ehs.ucr.edu/laboratory/hotplatesafetyadvisory20110715.pdf
http://safety.ucanr.edu/files/152250.pdf
https://www.drs.illinois.edu/News/Warning-about-malfunctioning-hotplates